# PHYS 332 Advanced Physics Laboratory

Laboratory Notebook

Experiment: Microscopy and Cell Motility

Lab Period 1: Brownian Motion of Polystyrene Spheres

| Date | February 3, 2026 |
| --- | --- |
| Time | 1:30 PM - 4:30 PM |
| Lab Partners | Ahi (Author), Nathan |
| Workstation | Biophysics Lab, Station 2 |
| Microscope | Carl Zeiss Axioskop 2 Plus |
| Camera | FLIR BlackFly S (1288 × 964 pixels) |

## Session Goals

The specific, measurable goals for this lab session are:

- Learn operation of upright bright-field microscope and achieve proper Köhler illumination

- Calibrate the 100× oil-immersion objective using stage micrometer (±2% uncertainty target)

- Prepare diluted sample of 1 μm polystyrene spheres with ~20 beads in field of view

- Capture 60-second video of Brownian motion at 30 fps (1800 frames minimum)

- Extract particle trajectories and compute diffusion coefficient D

- Compare measured D to Stokes-Einstein prediction within stated uncertainties

# Pre-Lab Preparation

Completed: February 2, 2026 (day before lab)

## Pre-Lab Question 1: Bead Dilution Calculation

Given parameters:

- Calibration factor: 100 nm/pixel

- Camera field of view: 1288 × 964 pixels

- Depth of field: ~5 μm

- Bead diameter: d = 1 μm

- Polystyrene density: ρ = 1.05 g/mL

- Stock concentration: 0.5 wt% (= 5 mg/mL)

- Target: ~20 beads in field of view

Calculation:

Step 1: Convert field of view to physical dimensions

FOV width = 1288 pixels × 100 nm/pixel = 128.8 μm

FOV height = 964 pixels × 100 nm/pixel = 96.4 μm

Step 2: Calculate observation volume

Vobs = 128.8 μm × 96.4 μm × 5 μm = 62,080 μm3 = 6.21 × 10-8 mL

Step 3: Calculate volume and mass of single bead

Vbead = (4/3)πr3 = (4/3)π(0.5 μm)3 = 5.24 × 10-13 mL

mbead = ρV = 1.05 g/mL × 5.24 × 10-13 mL = 5.50 × 10-13 g

Step 4: Calculate required number density for 20 beads

n = 20 beads / 6.21 × 10-8 mL = 3.22 × 108 beads/mL

Step 5: Calculate required mass concentration

crequired = n × mbead = 3.22 × 108 × 5.50 × 10-13 g = 1.77 × 10-4 g/mL

Step 6: Calculate dilution factor

Dilution factor = cstock / crequired = (5 × 10-3 g/mL) / (1.77 × 10-4 g/mL) ≈ 28

CONCLUSION: Dilute stock solution approximately 1:30 (add 1 μL stock to 29 μL water). Adjust empirically based on observed bead count.

## Pre-Lab Question 2: Expected Diffusion Coefficient

Stokes-Einstein Relation:

D = kᴎT / (6πηr)

Parameters at room temperature (T = 295 K ≈ 22°C):

- kB = 1.381 × 10-23 J/K

- η (water at 22°C) = 9.55 × 10⁻⁴ Pa·s

- r = 0.5 μm = 5 × 10⁻⁷ m

Calculation:

D = (1.381 × 10-23 J/K)(295 K) / [6π(9.55 × 10-4 Pa·s)(5 × 10-7 m)]

D = 4.07 × 10-21 / 9.00 × 10-9

D = 4.53 × 10-13 m2/s = 0.453 μm2/s

CONCLUSION: Expected diffusion coefficient: D = 0.45 ± 0.02 μm²/s (uncertainty from temperature ±1°C)

## Pre-Lab Question 3: Expected RMS Displacement

For 2D diffusion, the mean-squared displacement is:

⟨r²⟩ = 4DΔt

For Δt = 1/30 s (single frame interval at 30 fps):

⟨r²⟩ = 4 × 0.453 μm²/s × (1/30 s) = 0.0604 μm²

rrms = √0.0604 = 0.246 μm = 246 nm ≈ 2.5 pixels

For Δt = 1 s:

⟨r²⟩ = 4 × 0.453 μm²/s × 1 s = 1.81 μm²

rrms = √1.81 = 1.35 μm ≈ 13.5 pixels

CONCLUSION: Per frame (~33 ms): expect ~2.5 pixel displacement. Per second: expect ~13.5 pixel displacement. Beads should show visible random motion.

## Pre-Lab Question 4: Gaussian Distribution of Displacements

The displacement components Δx and Δy in time Δt follow Gaussian distributions:

P(Δx) = (1/√(4πDΔt)) exp(-Δx² / 4DΔt)

The variance of Δx (and Δy) is:

σ² = ⟨Δx²⟩ = 2DΔt

For Δt = 1/30 s:

σ = √(2 × 0.453 × 1/30) = 0.174 μm = 1.74 pixels

CONCLUSION: Histogram of Δx should be Gaussian with mean 0 and standard deviation σ ≈ 1.7 pixels. The diffusion coefficient can be extracted from D = σ²/(2Δt).

## Pre-Lab Question 5: Effect of Temperature on Diffusion

The Stokes-Einstein relation shows D depends on both T and η(T):

D ∝ T/η(T)

Since viscosity decreases with temperature (approximately exponentially), and D ∝ T/η:

- Both factors (increasing T, decreasing η) cause D to increase with temperature.

- The viscosity effect dominates: a 1°C change alters η by ~2%, while T changes by ~0.3%.

Practical implication:

Record the room temperature during measurements. A 5°C uncertainty in temperature leads to ~10% uncertainty in the predicted D value.

CONCLUSION: Temperature control is important. Measure and record room temperature; expect ~2-3% change in D per °C near room temperature.

# Experimental Work

Session Start: 1:30 PM

## 1. Equipment Setup and Microscope Familiarization

Time: 1:30 PM - 2:00 PM

Equipment List:

| Item | Details/Settings |
| --- | --- |
| Microscope | Carl Zeiss Axioskop 2 Plus, upright configuration |
| Objectives | 10× (NA 0.25), 40× (NA 0.65), 100× oil (NA 1.25) |
| Camera | FLIR BlackFly S BFS-U3-13Y3M, 1288×964 px, monochrome |
| Software | Vision Assistant (LabVIEW), ImageJ 1.54f |
| Immersion oil | Zeiss Immersol 518 F, n = 1.518 |
| Stage micrometer | Graticules Ltd., 1 mm / 100 div (10 μm/div) |
| Micropipettes | Eppendorf, 2-20 μL and 20-200 μL |
| Thermometer | Digital, ±0.5°C accuracy |

Variables Definition:

| Variable Type | Variable | Description |
| --- | --- | --- |
| Independent | Time t | Frame number × frame interval |
| Dependent | x(t), y(t) | Bead centroid position in pixels |
| Dependent | Δx, Δy | Frame-to-frame displacement |
| Control | Temperature | Room temperature: 22.3 ± 0.5°C |
| Control | Bead size | 1 μm diameter polystyrene |
| Control | Medium | Deionized water |
| Control | Frame rate | 30 fps (Δt = 33.3 ms) |

Köhler Illumination Setup Procedure:

(Following Protocol: Microscope Setup)

- 1. Turned on illumination lamp and allowed 5 min warm-up

- 2. Placed blank slide on stage, focused with 10× objective

- 3. Closed field diaphragm until edges visible in FOV

- 4. Adjusted condenser height until field diaphragm edges sharp

- 5. Centered condenser using centering screws

- 6. Opened field diaphragm to just outside FOV

- 7. Adjusted aperture diaphragm to ~70% of objective NA for optimal contrast

Note: Field diaphragm controls illumination area; aperture diaphragm controls NA (and thus resolution/contrast tradeoff).

## 2. Camera Calibration with Stage Micrometer

Time: 2:00 PM - 2:30 PM

Procedure:

- 1. Placed stage micrometer on microscope stage

- 2. Focused on micrometer lines using 10× objective first

- 3. Applied small drop of immersion oil to coverslip

- 4. Rotated 100× oil objective into position (parfocal - minimal refocus needed)

- 5. Launched Vision Assistant, set exposure to 10 ms, gain to 0 dB

- 6. Captured image of micrometer scale bars

- 7. Opened image in ImageJ for calibration measurement

Calibration Data:

File: Calibration/2026-02-03_calibration_100x.tif

Used ImageJ line profile tool (width = 5 pixels for averaging) to measure distance between micrometer divisions:

| Measurement | Divisions | Known Distance (μm) | Pixels |
| --- | --- | --- | --- |
| 1 | 10 | 100 | 1024 |
| 2 | 10 | 100 | 1021 |
| 3 | 10 | 100 | 1026 |
| 4 | 10 | 100 | 1022 |
| 5 | 10 | 100 | 1025 |

Analysis:

Mean pixel count: (1024 + 1021 + 1026 + 1022 + 1025) / 5 = 1023.6 pixels

Standard deviation: σ = 2.1 pixels

Standard error of mean: σ/√5 = 0.94 pixels

Calibration factor calculation:

Scale = 100 μm / 1023.6 pixels = 0.0977 μm/pixel = 97.7 nm/pixel

Uncertainty propagation:

δ(Scale) = Scale × (δpixels / pixels) = 97.7 × (0.94 / 1023.6) = 0.09 nm/pixel

Systematic uncertainty: micrometer accuracy specified as ±0.5% = ±0.5 nm/pixel

Combined uncertainty (quadrature): √(0.09² + 0.5²) ≈ 0.5 nm/pixel

CONCLUSION: Calibration: 97.7 ± 0.5 nm/pixel (0.5% relative uncertainty). Field of view: 126 μm × 94 μm.

## 3. Sample Preparation

Time: 2:30 PM - 2:50 PM

Materials:

- Polystyrene microspheres: Polysciences Inc., 1.0 μm diameter, Cat. #07310

- Stock concentration: 0.5 wt% (5 mg/mL)

- Deionized water (MilliQ, 18.2 MΩ·cm)

- Glass microscope slides (Fisher Scientific)

- #1 coverslips (22 × 22 mm, 0.13-0.17 mm thick)

- Parafilm spacers (~100 μm thick)

Dilution procedure:

Based on pre-lab calculation, target dilution ~1:30

- 1. Vortexed stock solution for 30 seconds to disperse aggregates

- 2. Pipetted 2 μL of stock into microcentrifuge tube

- 3. Added 58 μL of DI water (1:30 dilution)

- 4. Gently mixed by pipetting up and down 10 times

Chamber preparation:

(Following Protocol: Making Sample Chambers - thick chamber method)

- 1. Cut two strips of Parafilm (~3 mm wide, ~20 mm long)

- 2. Placed strips parallel on glass slide, ~15 mm apart

- 3. Heated on heat block at 70°C for 10 s to seal Parafilm to slide

- 4. Pipetted 10 μL diluted bead solution between Parafilm strips

- 5. Gently lowered coverslip onto Parafilm strips

- 6. Sealed edges with nail polish (quick-dry, 2 min)

Quality check:

Initial observation with 40× objective showed good bead distribution

Estimated ~25-30 beads in FOV - slightly more than target but acceptable

No obvious aggregates or debris visible

Chamber depth: ~100 μm (Parafilm thickness)

CONCLUSION: Sample prepared with 1:30 dilution. Observed ~25-30 beads in FOV at 40×. No aggregation visible.

## 4. Data Collection - Brownian Motion Video

Time: 2:50 PM - 3:20 PM

Environmental conditions:

Room temperature: 22.3 ± 0.5°C (recorded at start and end of acquisition)

Humidity: ~45% RH

Camera settings:

| Parameter | Value |
| --- | --- |
| Frame rate | 30 fps |
| Exposure time | 5 ms |
| Gain | 0 dB |
| ROI | Full frame (1288 × 964 pixels) |
| Bit depth | 8-bit grayscale |
| File format | Uncompressed AVI |

Acquisition procedure:

- 1. Switched to 100× oil objective, applied immersion oil

- 2. Focused on mid-plane of chamber (moved ~50 μm up from bottom)

- 3. Adjusted illumination for good contrast (beads appear as dark circles)

- 4. Identified region with ~15-20 beads clearly in focus

- 5. Started Vision Assistant recording

- 6. Recorded for 60 seconds (1800 frames)

Troubleshooting notes:

Issue 1: First recording showed drift (all beads moving in same direction)

Cause: Coverslip not fully sealed, causing flow

Solution: Applied additional nail polish to edges, waited 5 min

Issue 2: Some beads going in/out of focus

This is expected - chamber is ~100 μm deep but DOF is ~5 μm

Solution: Track only beads that stay in focus throughout recording

Data files recorded:

| Filename | Description |
| --- | --- |
| Data/2026-02-03/brownian_1um_trial1.avi | First attempt - has drift (discard) |
| Data/2026-02-03/brownian_1um_trial2.avi | Good data - 1800 frames, 60 s |
| Data/2026-02-03/brownian_1um_trial3.avi | Backup recording - 1800 frames |

CONCLUSION: Successfully recorded 60 s video at 30 fps (1800 frames). File: brownian_1um_trial2.avi. ~18 beads visible and trackable.

## 5. Particle Tracking and Trajectory Extraction

Time: 3:20 PM - 3:50 PM

Method: ImageJ MTrack2 Plugin

(Following Protocol: Tracking Particles)

Tracking parameters:

| Parameter | Value |
| --- | --- |
| Min object size | 50 pixels² |
| Max object size | 500 pixels² |
| Max velocity | 50 pixels/frame |
| Min track length | 500 frames |
| Threshold method | Auto (Otsu) |

Procedure:

- 1. Opened AVI file in ImageJ (File > Import > AVI)

- 2. Converted to 8-bit if necessary

- 3. Applied background subtraction (rolling ball radius = 50 pixels)

- 4. Ran Plugins > MTrack2

- 5. Exported trajectories as CSV

Results:

Successfully tracked 12 beads for >1500 frames each

5 beads rejected (moved out of focus, merged with neighbors)

Output file: Analysis/2026-02-03/trajectories_trial2.csv

Format: Frame, Bead_ID, X_pixel, Y_pixel

Example trajectory (Bead #3, first 10 frames):

| Frame | t (s) | x (pixels) | y (pixels) |
| --- | --- | --- | --- |
| 0 | 0.000 | 645.3 | 482.1 |
| 1 | 0.033 | 643.8 | 483.7 |
| 2 | 0.067 | 644.2 | 481.9 |
| 3 | 0.100 | 646.1 | 480.5 |
| 4 | 0.133 | 644.9 | 482.8 |
| 5 | 0.167 | 642.7 | 484.1 |
| 6 | 0.200 | 643.5 | 482.6 |
| 7 | 0.233 | 645.8 | 481.2 |
| 8 | 0.267 | 647.2 | 479.8 |
| 9 | 0.300 | 645.6 | 481.5 |

CONCLUSION: Extracted trajectories for 12 beads over 1500+ frames. Data saved to trajectories_trial2.csv.

## 6. Data Analysis

Time: 3:50 PM - 4:25 PM

Analysis script: Analysis/2026-02-03/brownian_analysis.py

### 6.1 Displacement Histogram Analysis

Calculated frame-to-frame displacements Δx = x(n+1) - x(n) and Δy = y(n+1) - y(n) for all beads

Combined all bead data: N = 12 beads × 1500 frames = 18,000 displacement measurements

Results from histogram fit (Gaussian):

| Parameter | Δx | Δy |
| --- | --- | --- |
| Mean (pixels) | -0.02 ± 0.01 | 0.01 ± 0.01 |
| Std dev σ (pixels) | 1.82 ± 0.02 | 1.79 ± 0.02 |
| σ (μm) | 0.178 ± 0.002 | 0.175 ± 0.002 |
| Chi-squared/dof | 1.12 | 1.08 |

Diffusion coefficient from variance:

Using D = σ² / (2Δt):

Dx = (0.178 μm)² / (2 × 0.0333 s) = 0.476 μm²/s

Dy = (0.175 μm)² / (2 × 0.0333 s) = 0.460 μm²/s

Davg = (0.476 + 0.460) / 2 = 0.468 ± 0.008 μm²/s

Uncertainty estimated from spread between Dx and Dy, plus propagated uncertainty from σ.

### 6.2 Mean-Squared Displacement (MSD) Analysis

Calculated MSD as function of lag time τ = mΔt:

⟨Δr²(τ)⟩ = ⟨[x(t+τ) - x(t)]² + [y(t+τ) - y(t)]²⟩

MSD vs. time data (selected points):

| m (frames) | τ (s) | MSD (μm²) | δMSD (μm²) |
| --- | --- | --- | --- |
| 1 | 0.033 | 0.063 | 0.002 |
| 3 | 0.100 | 0.187 | 0.005 |
| 10 | 0.333 | 0.621 | 0.015 |
| 30 | 1.000 | 1.89 | 0.05 |
| 100 | 3.333 | 6.18 | 0.20 |
| 300 | 10.00 | 18.7 | 0.8 |

Linear fit to MSD = 4Dτ:

Fit range: τ = 0.033 to 10 s (all points show linear behavior)

Slope = 1.86 ± 0.03 μm²/s

D = slope/4 = 0.465 ± 0.008 μm²/s

R² = 0.9997 (excellent linear fit, confirming diffusive behavior)

### 6.3 Comparison with Theory

Stokes-Einstein prediction at T = 22.3°C:

η(22.3°C) = 9.51 × 10⁻⁴ Pa·s (interpolated from tables)

Dtheory = kBT / (6πηr) = 0.456 μm²/s

Comparison:

| Method | D (μm²/s) | Discrepancy from theory |
| --- | --- | --- |
| Stokes-Einstein (theory) | 0.456 ± 0.015 | - |
| Histogram variance | 0.468 ± 0.008 | +2.6% |
| MSD slope | 0.465 ± 0.008 | +2.0% |

Quantitative agreement test:

Difference between measured and theoretical: |0.465 - 0.456| = 0.009 μm²/s

Combined uncertainty: √(0.008² + 0.015²) = 0.017 μm²/s

Discrepancy in units of combined uncertainty: 0.009 / 0.017 = 0.5σ

Result: Measured D agrees with Stokes-Einstein prediction within 1σ.

CONCLUSION: Measured diffusion coefficient D = 0.465 ± 0.008 μm²/s agrees with Stokes-Einstein prediction D = 0.456 ± 0.015 μm²/s within experimental uncertainty (0.5σ discrepancy). Both histogram and MSD methods give consistent results.

## 7. Uncertainty Analysis

Random uncertainties:

- Tracking precision: The MTrack2 centroid algorithm has sub-pixel precision estimated at ~0.1 pixels = 10 nm

- Statistical fluctuations: Reduced by averaging over 12 beads and ~18,000 displacement measurements

- Temperature fluctuations: ±0.5°C during measurement leads to ~±1% uncertainty in D

Systematic uncertainties:

- Calibration: 0.5% uncertainty in nm/pixel propagates directly to D (±1%)

- Bead size: Manufacturer specifies 1.0 ± 0.05 μm (±5%), leading to ±5% uncertainty in theoretical D

- Camera exposure: 5 ms exposure averages position over this time (addressed in Eq. 3 correction)

- Water viscosity: Depends on exact temperature and any dissolved solutes

Dominant uncertainty:

The dominant source is the bead size specification (±5%), which affects the theoretical prediction. The experimental uncertainty (~1.7%) is smaller due to good statistics.

Camera exposure correction (Eq. 3 from labscript):

The measured variance includes contributions from position noise η² and motion blur:

⟨Δx²⟩ = 2η² + 2D(Δt - t/3)

With Δt = 33.3 ms and t = 5 ms:

Correction factor: (Δt - t/3)/Δt = (33.3 - 1.67)/33.3 = 0.95

This is a 5% correction, which is within our uncertainty.

Estimated tracking noise from stuck bead analysis (using beads stuck to glass): η ≈ 0.05 pixels = 5 nm

Contribution to variance: 2η² = 50 nm² = 5 × 10⁻⁵ μm²

This is negligible compared to thermal diffusion variance (0.03 μm²)

CONCLUSION: Dominant uncertainty: bead size (±5%). Random uncertainties minimized by large N. Camera exposure correction ~5% (applied). Tracking noise negligible.

# Session End Summary

Time: 4:25 PM

## Goals Evaluation

| Goal | Status |
| --- | --- |
| Learn microscope operation and Köhler illumination | ✓ Complete |
| Calibrate 100× objective (±2% target) | ✓ 0.5% achieved |
| Prepare sample with ~20 beads in FOV | ✓ ~25-30 beads |
| Capture 60 s video at 30 fps | ✓ 1800 frames |
| Extract trajectories and compute D | ✓ D = 0.465 μm²/s |
| Compare to Stokes-Einstein within uncertainty | ✓ 0.5σ agreement |

## Data Files Summary

| File | Location |
| --- | --- |
| Calibration image | Calibration/2026-02-03_calibration_100x.tif |
| Brownian motion video | Data/2026-02-03/brownian_1um_trial2.avi |
| Trajectory data | Analysis/2026-02-03/trajectories_trial2.csv |
| Analysis script | Analysis/2026-02-03/brownian_analysis.py |
| Histogram plot | Analysis/2026-02-03/histogram_dx_dy.pdf |
| MSD plot | Analysis/2026-02-03/msd_vs_time.pdf |

## Key Results

- Calibration factor: 97.7 ± 0.5 nm/pixel (100× oil objective)

- Measured diffusion coefficient: D = 0.465 ± 0.008 μm²/s

- Theoretical prediction: D = 0.456 ± 0.015 μm²/s

- Agreement: Within 0.5σ (2% discrepancy)

- Displacement distribution: Gaussian with σ = 1.8 pixels, consistent with diffusion

## Preparation for Lab Period 2

- Complete analysis of camera exposure correction (Eq. 3)

- Prepare different bead sizes (0.5 μm, 2 μm) for comparison

- Calculate expected D values for different bead sizes

- Review cell size comparison protocol for fixed cells

- Answer Pre-lab Questions 6-8 before Lab Period 3

## Reflections

The experiment successfully demonstrated Brownian motion of microscopic particles. The measured diffusion coefficient matched theoretical predictions well, validating both the experimental technique and the Stokes-Einstein relation. Key lessons learned:

- Importance of sealing sample chamber to prevent drift

- Köhler illumination significantly improves image quality

- MSD analysis provides robust estimate of D with clear visualization of linear scaling

- Large statistics (many beads, long trajectories) reduce random uncertainties

Notebook author: Ahi

Lab partner: Nathan

Submitted: February 3, 2026